## Slide 1
Electro-optic qualification of LSST production rafts P. O’Connor1*, Andrei Nomerotski1, Homer Neal2 1Brookhaven National Laboratory, Upton, NY 2SLAC National Laboratory, Menlo Park, CA * Subsystem scientist, LSST science rafts Scientific Detector Workshop 2017

## Slide 2: Brookhaven Instrumentation detector development: first 35 years
Silicon drift detector 1983
Cylindrical SDD for x-ray spectroscopy
Preamp optimal filtering1983

## Slide 3: Last 25 years
640-element strip detector
HERMES readout
high resolution hexagonal SDD array
SDD readout ASIC
8e- noise
j
32-channel RatCAP ASIC
Compact Gamma Camera
image of conscious rat brain
384-element EXAFS array
HERMES readout
silicon vertex tracker
240-channel multichp module
time expansion chamber
24-channel ASIC preamp/shaper
Electromagnetic calorimeter
sampling/digitizing board
30,000 channels total
quad hybrid preamp
TEC front end module
16-element SDD array

## Slide 4: A large detector

## Slide 5: Observatory construction at Cerro Pachon
3.2Gpixel camera on 8.4m telescope
Dedicated 10-year deep optical-NIR survey
Rapid cadence to image entire sky to r=24 every 3 nights
Broad science goals:
  Cosmology: dark matter and dark energy
  Milky Way assembly history
  Solar System minor bodies
  Transients

## Slide 6: Telescope, camera, and focal plane
3024 video channels
4347 clock and bias signals
2 sec frame readout
  550kpix/s pixel rate
< 9e- rms noise
18b digitization
Focal plane thermal control
System health monitoring
Sensor protection
1200 W total power
  400 mW/channel
170 l volume
Vacuum operation
Flat ±10um
requirements

## Slide 7: Raft Tower Module
Photons in – bits out
Complete 12K x 12K imager
Support sensors mechanically to meet strict coplanarity
Thermal management of sensors and electronics
Provide bias, timing, and control signals for CCD operation
144 video channels with low noise analog signal processing, digitizing, multiplexing of pixel data
Diagnostics + monitoring
12.7 x 12.7 x 42cm3
20 kg
39W

## Slide 8: Detectors
100mm-thick, high resistivity silicon CCDs, fully-depleted with transparent conductive window and AR coating, back-illuminated
  for broadband QE and small PSF
4K x 4K format
  4 die sites/6”wafer
16-fold parallel output
  for low noise readout at 2s frame read time
10mm pixels
  for optimum sampling at LSST plate scale
Buttable, thermally-matched packaging
  >92% fill factor
Flatness and alignment tolerance to bring image surface of all CCDs coplanar to ±9mm
  for use in fast f/1.2 beam
Mechanical mounting and alignment features and electrical interface specified
  Interchangeability between vendors
Operating temperature -100C
  Suppress dark current to negligible levels

## Slide 9: 2 detector variants
Uniform illumination
Test target image
CCD in test fixture
T-E2v CCD250
T-DALSA/ITL STA3800C
(overscan displayed)
(overscan suppressed)

## Slide 10: Raft electronics
Functions:
Video CDS processing
18-bit digitization and serialization
HV clock and bias generation
CCD temperature monitoring and control
Onboard diagnostic features including waveform recorder
Implementation:
Three 4” x 16”, 14-layer boards
Video processing in CMOS ASIC (x18)
Xilinx XC7K160T FPGA (x3)
AD7982 18-bit 1MSa/s SAR ADC (x144)
+40/±15/+8/+7/+5V supplies, onboard regulators

## Slide 11
functional test
Data base
ASIC
REB
functional test
Data base
Burn-in,
Thermal cycle

## Slide 12: Raft tower assembly (1)

## Slide 13: Raft tower assembly (2)

## Slide 14: Completed Raft Tower

## Slide 15: Warm connectivity test
0 sec image
4 sec image
SF bias currents
counts vs. channel

## Slide 16: Tower test cryostat

## Slide 17: Cryocooler refrigeration capacities
Single PT-30 cryo head for RSA
Dual NF-55 cold head for electronics
2X
DT
expected Q
DT
expected Q
PT-30

## Slide 18: Tower in cryostat
Friday, Jan 27 2017
ETU#1 in TS-8 at BNL

## Slide 19: RTM production status Sept. 2017
ETU1
ETU2
RTM8
RTM1
RTM2
RTM3
RTM4
RTM5
RTM6
RTM7
RTM9
RTM10
RTM11
RTM22
…
T-e2v sensors
ITL sensors
future production
assembled and tested
this presentation
we’re going to need a bigger assembly bench…

## Slide 20: Bias images: 3 towers (432 channels)
Bias10%FS
max FWC 82%FS

## Slide 21: First step – gain determination
144 channel 55Fe spectra
144 channel PTC

## Slide 22: 350nm flatfields

## Slide 23: 750nm flatfields
4389

## Slide 24: 1000nm flatfields
4389

## Slide 25: Intra- and inter-raft uniformity
Signal vs. exp time 144 chan. RTM2
QE vs. l 144 chan. RTM2
Mean intra-CCD PRNU
Mean inter-CCD PRNU

## Slide 26: EO performance statistics: 3 rafts
gain e-/ADU | 
spec | --
<x> | 0.70
1s | 0.02
95% | 0.08
read noise e- | 
spec | 9
<x> | 4.66
1s | 0.16
95% | 0.68
full well ke- | 
spec | <175
<x> | 136
1s | 5.5
95% | 20.3
PSF s | 
spec | 5
<x> | 4.12
1s | .062
95% | .160
nonlin % | 
spec | 3
<x> | 0.15
1s | 0.03
95% | 0.11
HCTI ppm | 
spec | 5
<x> | 0.72
1s | 0.76
95% | 2.69

## Slide 27: Parameter maps
gain e-/ADU
noise e-
nonlinearity %
PTC_gain e-/ADU
full well e-
dark current e-/pix/s

## Slide 28: Within-CCD crosstalk
victim vs. aggressor pixel
aggressor in segment 6
Full 16x16 crosstalk matrix (prototype elx.)
Xtalk from single aggressor

## Slide 29: Noise correlations – good, bad, ugly

## Slide 30: Performance goals
  2 second readout time
  9e- read noise
  QE > 75% (450 < l < 870nm)
  QE > 20% from 320 < l < 1050nm
  > 99.9995 % CTE serial and parallel
  99.0% operable pixels
  91.4% area fill factor
  < 0.24’’ FWHM diffusion PSF
  0.3% electronic crosstalk
  < 58W power per RTM

## Slide 31: Performance goals achieved
  2 second readout time 1.975s
  9e- read noise 4.7 ± 0.17 e-
  QE > 75% (450 < l < 870nm) 85%
  QE > 20% from 320 < l < 1050nm 23%
  > 99.9995 % CTE serial and parallel 99.9999%
  99.0% operable pixels 99.97%
  91.4% area fill factor 92.97%
  < 0.24’’ FWHM diffusion PSF 0.18”
  0.3% electronic crosstalk 0.03%
  < 58W power per RTM 39W

## Slide 32
THANK YOU

## Slide 33: REFERENCES
SDW2013 “System Tests of LSST Integrated Readout Chain” Florence, POC
SPIE2016 “Integrated System Tests of the LSST Raft Tower Modules”, Edinburgh, POC
”Electro-optical testing of fully depleted CCD image sensors for the Large Synoptic Survey Telescope camera”, Vol. 9154 (2014), pp. 915418-915418-17, doi:10.1117/12.2056733 by Peter E. Doherty, Pierre Antilogus, Pierre Astier, et al.
LSST sensor requirements and characterization of the prototype LSST CCDsV Radeka, J Frank, JC Geary, DK Gilmore, I Kotov, P O'Connor, P Takacs, Journal of Instrumentation 4 (03), P03002
Development of the LSST raft tower modules, Society of Photo-Optical Instrumentation Engineers (SPIE) Conference Series, Vol. 8453 (July 2012), doi:10.1117/12.926645 by P. O'Connor, I. Kotov, P. Z. Takacs, et al.
Technology of the LSST focal plane, Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment, Vol. 582, No. 3. (01 December 2007), pp. 902-909, doi:10.1016/j.nima.2007.07.120, by P. Oconnor, J. Geary, K. Gilmore, J. Oliver, V. Radeka, P. Takacs
Lu, W., et al. "CCD emulator design for LSST camera." SPIE Astronomical Telescopes+ Instrumentation. International Society for Optics and Photonics, 2016.
O'Connor, P., J. Fried, and I. Kotov. "An advanced CCD emulator with 32MB image memory." SPIE Astronomical Telescopes+ Instrumentation. International Society for Optics and Photonics, 2012
P. O'Connor, I. Kotov, P. Z. Takacs, J. S. Frank, S. Plate, R. Van Berg, M. Newcomer, P. Antilogus, Hervé Lebbolo, V. Tocut, C. Juramy, P. Doherty, N. Felt, "Development of the LSST raft tower modules", Proc. SPIE 8453, High Energy, Optical, and Infrared Detectors for Astronomy V, 84530L (25 September 2012); doi: 10.1117/12.926645; http://dx.doi.org/10.1117/12.926645

## Slide 34: Cryocooler refrigeration capacities
Single PT-30 cryo head for RSA
Dual NF-55 cold head for electronics
2X
DT
expected Q
DT
expected Q

## Slide 35: Camera Control System (CCS) monitoring GUI

## Slide 38: Production status Sept. 2017
RTMs assembled: 10
Engineering models: 2
I think we’re going to need a bigger table…