Maintenance Procedures

Background

All gases other than helium can freeze in the cryopump refrigerator. During

manufacturing of Cryopump systems, gaseous impurities are removed using stringent

manufacturing control. The delivered system contains sufficiently low concentrations

of gaseous impurities so they are not of concern.

It is possible that additional gaseous contaminants can be introduced into the helium

circuit. These contaminates will collect in the cryopump refrigerator as frozen gas.

When enough frozen contaminates accumulate in the displacer, the amount of torque

required to drive the displacer mechanism will be greater then the amount of motor

torque available. This will result in the cryopump refrigerator motor slowing or

stalling.

Regardless of the refrigerator being used, small amounts of these gaseous

contaminants can be removed from the entire helium system by freezing them in the

Decontamination Coldhead Decontamination Unit. If there is gross contamination in

the helium circuit, then the entire system will need to be decontaminated and the

compressor adsorbers changed. Refer to the Alternate Decontamination Procedure

using Dilution Purging Techniques in this document. The use of this

decontamination procedure will return most Cryopumps to proper operation.

Helium Circuit Decontamination

Contamination of the helium-gas circuit is indicated by sluggish or

intermittent operation (ratchetting) of the cold head drive mechanism. With

severe contamination the cold head drive may seize and fail to operate. One

of the major sources of contamination is using helium gas of less than the

required purity. When performing the decontamination process, use only

99.999% pure-helium gas, and the regulator and charging line must be

properly connected and purged.

This decontamination procedure will remove contaminants from the cold

head and/or compressor, thereby restoring system performance. The coldtrapping

of contaminants inside the cold head during this procedure will

also decontaminate the compressor if the contamination of the system is

not severe. Separate decontamination of the compressor is required

whenever the compressor has been opened to atmosphere, or the pressure

dropped to zero.

Cryo-Torr Cryopump Decontamination Procedures

NOTE: Refer also to the appropriate Compressor Manual.

1. Cool down the cryopump and operate it for one to three hours.

If the system will not cool down, proceed to step 2. Operating

the cryopump will isolate the contaminants by "freezing" them

in the cold head. The contaminants in the helium-gas circuit of

the cryopump tend to become frozen inside the cold head. The

longer the cryopump is operated beyond the one-hour period,

the greater is the amount of contamination that becomes

isolated inside the cold head.

2. Shut down the cryopump.

3. Immediately disconnect the helium supply and return lines from the

gas-supply and gas-return connectors at the rear of the compressor.

Leave them attached to the cold head.

4. Attach the maintenance manifold to the disconnected ends of the

return and supply lines.

5. Reduce the pressure in the cold head to a level of 30 psig by using

the maintenance manifold. Reducing the pressure in the cold head

to below 30 psig (200 kPa) may introduce more contaminants into

the helium circuit.

If you only have the manual regeneration option, turn the cryopump

OFF and open the purge valve until the second stage temperature

reaches 290K.

6. Attach a two-stage regulator (0-3000/0-400 psig) and charging line

to a helium bottle (99.999% pure). DO NOT OPEN THE BOTTLE

VALVE AT THIS TIME. Purge the regulator and charging line as

introduced in steps a through d below. Do not use helium gas that is

less than 99.999% pure.

a. Open the regulator a small amount by turning the adjusting

knob clockwise until it contacts the diaphragm; then, turn

approximately 1/8 to 1/4 turn more, so that the regulator is

barely open.

b. Slowly open the bottle valve, and purge the regulator and

line for 10 to 15 seconds. Turn the regulator knob counterclockwise

until the helium stops flowing.

c. Loosely connect the charge line to the 1/8-inch Hoke valve

on the maintenance manifold.

d. Purge the charge line again, as in step a, for 30 seconds, and

tighten the charge line flare fitting onto the Hoke valve

while the helium is flowing.

This procedure is required to ensure that both the regulator and the

charging line will be purged of air and that the air trapped in the

regulator will not diffuse back into the helium bottle. For best results,

CTI-CRYOGENICS suggests a dedicated helium bottle, regulator, and

line, which are never separated, for adding helium.

7. Perform in sequence:

a. Backfill the cold head with helium to a static charge pressure

of 195-205 psig (1345-1415 kPa) by adjusting the regulator

to the required pressure, and opening the Hoke valve

on the manifold. Close the Hoke valve when the pressure is

correct.

b. Depressurize the cold head to 30 and 50 psig 200 and 330

kPa) by slowly opening the ball valve and allowing the

helium to bleed out slowly. Do not reduce the pressure to

less than 30 psig or the cold head may be further contaminated.

c. Perform the flushing steps a and b four more times.

d. Pressurize the cold head to a static charge pressure of 195-

205 psig (1345-1415 kPa) and run the cold head drive motor

for 10 to 30 seconds.

e. Perform steps b through d four more times for a total of 25

flushes and 5 drive-motor runs.

8. Verify that the cold head is pressurized to the static charge pressure

of 195-205 psig (1345-1415 kPa).

9. Disconnect the maintenance manifold from the helium return and

supply lines.

10. Reconnect the helium return and supply lines to the return and supply

connectors at the rear of the compressor. The cryopump is now

ready for operation.